**Contract Award Notice**

**REPUBLIC OF ARMENIA**

**DISEASE PREVENTION AND CONTROL PROJECT**

Credit No.: 5222

Contract Title: **Medical, diagnostic equipment, Medical supplies and Medical furniture for "Sevan MC’’ and “Artashat MC” CJSCs,comprising of 4 lots**

Lot 1- Supply, Installation and training of Diagnostic and Imaging equipment for “Sevan MC” and “Artashat MC” CJSC

Lot 2- Supply, Installation and training of Medical Laboratory equipment for “Sevan MC” and “Artashat MC” CJSC

Lot 3 - Supply, Installation and training of Surgery and Hospital equipment for “Sevan MC” and “Artashat MC” CJSC

Lot4- Supply and Installation of Medical equipment and medical Furniture for “Sevan MC” and “Artashat MC” CJSC

**Reference No**.: CR4/ICB/B-G/018-17 (/1;2;3;4)

1. **Name of each Bidder who submitted a Bid; Bid prices as read out and evaluated prices:**

For Lot 1 are:

|  |  |  |
| --- | --- | --- |
| **Name of each Bidder** | **Bid prices as read out** | **Evaluated prices** |
|  “Ar. Medtechnica” LLC  | 214,090,600 | - |
| “Telia-Med” CJSC  | 226,906,480 | - |
| “A-Medeq” LLC P.O. | EUR 571,857  | - |
| “Concern-Energomash” CJSC  |  251,052,000 | 251,052,000 |

For Lot 2 are:

|  |  |  |
| --- | --- | --- |
| **Name of each Bidder** | **Bid prices as read out** | **Evaluated prices** |
|  “Telia-Med” CJSC | 43,040,090 | - |
| “Pink Flamingo B. & H. Center” LTD | 39,999,000 | 37,551,250 |
| “Roma” LLC |  47,361,000 | 47,361,000 |

For Lot 3 are:

|  |  |  |
| --- | --- | --- |
| **Name of each Bidder** | **Bid prices as read out** | **Evaluated prices** |
| “Ar. Medtechnica” LLC | 288,135,600 | 288,135,600 |
| “A-Medeq” LLC P.O. | EUR 491,021.2 | 289,486,341 |
| “Concern-Energomash” CJSC | 281,638,300 |  - |
| “Medtechservice” LLC | 182,521,700 | - |

For Lot 4 are:

|  |  |  |
| --- | --- | --- |
| **Name of each Bidder** | **Bid prices as read out** | **Evaluated prices** |
| “Concern-Energomash” CJSC & “T-Industry” CJSC JV | 198,255,700 | 198,255,700 |
| “Agat-777” LLC  | 207,109,700 | 207,109,700 |
| “Zorashen” LLC  | 221,351,200 | 221,351,200 |
| “Medimpex” LLC  |  215,018,080 | - |

1. **Name of bidders whose bids were rejected and the reasons for their rejection; and**

***Lot 1-* Supply, Installation and training of Diagnostic and Imaging equipment for “Sevan MC” and “Artashat MC” CJSC**

Submitted bids from “Ar. Medtechnica” LLC and “Telia-Med” CJSC were rejected, since there were considered as substantially non-responsive and not acceptable for detailed examination, as per technical evaluation bids have number of major deviations from the required technical aspects. Especially, following major deviations were identified:

1. “Ar. Medtechnica” LLC-
* Item 5. Universal X-Ray Unit: From required parameter for Table tilt motor driven, Anode heat storage capacity, Anodes drive, Nominal output of anode (Small and large), Tube column longitudinal excursion, Table film distance, tomography function.
1. “Telia-Med” CJSC-
* For Item #5 Universal X-Ray Unit: From required parameter for Table tilt motor driven, height of table, large focus, Tube Assembly Swivel, Tabletop-film distance, Computed/Digital Radiography system.
* For Item #7 Portable ultrasound (3 probes) from required parameter of image zoom, tissue specification imaging, sample volume and position, velocity range and time of battery operating.
1. “A-Medeq” LLC P. O.-Bid was rejected, since the company did not provide detailed comprising table of technical specifications of the suggested equipment (requested acc. To ITB 11.1 /j/ and ITB 16.3, on 26 January 2018). The bidder did not provide clarifications of its bid by the date and time set in the Purchaser’s request for clarification. Also the company was notified about conditions of ITB 27.2, additionally.

***Lot 2-* Supply, Installation and training of Medical Laboratory equipment for “Sevan MC” and “Artashat MC” CJSC**

Submitted Bid from “Telia-Med” CJSC; was rejected, since it was considered as substantially non-responsive and not acceptable for detailed examination, as per technical evaluation bid have number of major deviations from the required technical aspects. Especially, following major deviations were identified:

1. “Telia-Med” CJSC-
* Item 1. Hematological analyzer–From required parameters for fully automated parameter of Hematological analyzer and required accessories.
* Item 2. Semi-Automatic Chemistry Analyzer- From required parameters for photometric range and Flowcell.
* Item 3. Coagulometer- From required parameters for performed tesets.
* Item 4. Critical Care Analyser odel Convergys liquical BG Analyzer (BG + ISE + M + Hb), From required parameters for acceptable samples.

***Lot 3 -* Supply, Installation and training of Surgery and Hospital equipment for “Sevan MC” and “Artashat MC” CJSC**

Submitted bids from “Medtechservice” LLC and “Concern-Energomash” CJSC were rejected, since there were considered as substantially non-responsive and not acceptable for detailed examination, as per technical evaluation bids have number of major deviations from the required technical aspects. Especially, following major deviations were identified:

1. “Medtechservice” LLC-
* Item 4. Examination Lamp - From required parameter beam angle.
* Item 5. Cleaner–Recirculators - From required function for bactericidal effectivity.
* Item 6. Syringe pump - From required parameters for Flow rate increment and KVO.
* Item 7. Volumetric Infusion Pump - From required micro-drip function, parameters for Occlusion pressure.
* Item 17. Defibrillator with monitor - From required accumlator capacity.
* Item 18. Oxygen Concentrator - From required USB interface availability.
* Item 20. Suction pump - From required flow rate.
* Item 22. Videogastroscope and video colonofiberoscope - From required functions of Light source and Video processor.
* Item 27. Ventilator (Intensive ventilation equipment) - From required avalability of ADULTS AND PEDIATRICS modes, parameters for Programmable and Resulting Minute Volume, Support Pressure, Inspiratory Flow, Continuous Flow and Limited Pressure.
1. “Concern-Energomash” CJSC-
* Item 6. Syringe pump - From required parameters for Flow rate increment.
* Item 7. Volumetric Infusion Pump - From required micro-drip function and Occlusion pressure.
* Item 17. Defibrillator with monitor - From required parameters of energy steps, ECG monitoring and accumlator capacity.
* Item 22. Videogastroscope and video colonofiberoscope - From required functions of Video processor.
* Item 27. Ventilator (Intensive ventilation equipment) - From required avalability of ADULTS AND PEDIATRICS modes, parameter of respirator rate for adult, Inspiratory sensitivity, Limited Pressure and O2 consumption display.

***Lot4-* Supply and Installation of Medical equipment and medical Furniture for “Sevan MC” and “Artashat MC” CJSC**

Submitted bid from “Medimpex” LLC was rejected, since the company failed to list and price all items separatly in the Price Shedoule, as was reqeuested acc. to ITB 14.2 and ITB 14.6 (Prices quoted for each lot (contract) shall correspond at least to [100] percent of the items specified for each lot (contract).

**Name of the successful Bidder, and the Price it offered (evaluated), as well as the duration and summary scope of the contract awarded.**

Based on the result of evaluation (commercial and technical) the evaluation committee concluded to:recommend the award of the contract for:

|  |  |  |
| --- | --- | --- |
| **Name of the successful Bidder** | **Price it offered (evaluated)**  | **Duration and summary scope of the contract awarded** |
| **“Concern-Energomash” CJSC** | 251,052,000 | Supply and Installation of Diagnostic and Imaging equipment for "Sevan’’ and “Artashat” MCs  -Lot-1Delivery day-90 calendar daysInstallation of all equipment-10 daysTraining-5 days |
| **“Pink Flamingo B. & H. Center” LTD** | 37,551,250 | Supply and Installation of Laboratory equipment for "Sevan’’ and “Artashat” MCs -Lot-2 Delivery day-90 daysInstallation of all equipment-10 daysTraining-5 days |
| **“Ar. Medtechnica” LLC** | 288,135,600 | Supply, Installation and training of Surgery and Hospital equipment for "Sevan’’ and “Artashat” MCs Lot-3Delivery day-90 daysInstallation of all equipment-10 daysTraining-5 days |
| **“Concern-Energomash” CJSC & “T-Industry” CJSC JV** | 198,255,700 | Supply and Installation of Medical equipment and medical Furniture for "Sevan’’ and “Artashat” MCs -Lot-4Delivery day-90 daysInstallation of all goods-10 days |